## CB3012Plus (hidden)
CB3012 plus
Standard Function
Operating System : 2 Crank operation for High-Low and Head Rest positions
Position : - High – low: Manual operation with Crank
- Head rest: Manual operation with Crank,tilting from 0* to 80*
Primary : - Simple design, no sharp corners (smooth surface)
- Easy operating mechanism, with Crank System
- Sturdy and anti-rust
- Wheels that are easy to move and position,
with lock - unlock system
Lock System : Double Lock System on the Wheels
Accessories : IV Pole (2 Type Option) and Side Rail / Folding Side rail
Bed Side Cabinet, Over Bed Table, and Foot Step
Airmate Cpro Mattress / Urethane
Product Dimension
Dimension
L x W x D : 2230 x 1030 x 760~1005
Product Weight
Weight : ± 90KG
Load Capacity : ± 400KG (Static)
& ±170KG (Dynamic)
Packing Dimension : 2050 x 1070 mm x 190
MANUFACTURER LABEL
MANUFACTURE LABEL AND LOT NUMBER
SNI LABEL
IZIN EDAR KEMENKES
bahan dasar : sticker warna putih
posisi penempatan pada Frame

## CB3003Plus (hidden)
CB3003 plus
Standard Function
Operating System : 3 Crank operation for High-Low, Head Rest & Foot Rest positions
Position : - High – low: Manual operation with Crank
- Head rest: Manual operation with Crank,tilting from 0 to 80*
- Foot rest : manual operation with Crank, tilting form 0* to 80*
Primary : - Simple design, no sharp corners (smooth surface)
- Easy operating mechanism, with Crank System
- Sturdy and anti-rust
- Wheels that are easy to move and position,
with lock - unlock system
Lock System : Double Lock System on the Wheels
Accessories : IV Pole (2 Type Option) and Side Rail / Folding Side rail
Bed Side Cabinet, Over Bed Table, and Foot Step
Airmate Cpro Mattress / Urethane
Product Dimension
Dimension
L x W x D : 2230 x 1030 x 760~1005
Product Weight
Weight : ± 90KG
Load Capacity : ± 400KG (Static)
& ±170KG (Dynamic)
Packing Dimension : 2050 x 1070 mm x 190
MANUFACTURER LABEL
MANUFACTURE LABEL AND LOT NUMBER
SNI LABEL
IZIN EDAR KEMENKES
bahan dasar : sticker warna putih
posisi penempatan pada Frame

## CB3300Plus (hidden)
CB3300Plus
Standard Function
Operating System : Motor Actuator for High-Low, Foot Rest, and Head Rest positions
Position : - High – low: Automatic operation via handheld Controller
with Motor Actuator
- Head & Foot rest: Automatic operation via handheld controller
with Motor actuator
Primary : - Simple design, no sharp corners (smooth surface)
- Easy operating mechanism, with Motor Actuator
- Sturdy and anti-rust
- Wheels that are easy to move and position,
with lock - unlock system
Lock System : Total Lock System
Accessories : IV Pole (2 Type Option) and Side Rail / Folding Side rail
Bed Side Cabinet, Over Bed Table, and Foot Step
Airmate Cpro Mattress / Urethane
Product Dimension
Dimension
L x W x D : 2230 x 1030 x 760~1005
Product Weight
Weight : ± 90KG
Load Capacity : ± 400KG (Static)
& ±170KG (Dynamic)
Packing Dimension : 2050 x 1070 mm x 190
IZIN EDAR KEMENKES
bahan dasar : sticker warna putih
posisi penempatan pada Frame
MANUFACTURER LABEL
SNI LABEL
MANUFACTURE LABEL AND LOT NUMBER

## CB3012Plus petunjuk (hidden)
PETUNJUK PENGGUNAAN, MATERI PELATIHAN, PETUNJUK PEMASANGAN SERTA PEMELIHARAAN
No | TYPE | PENGGUNAAN | PEMASANGAN | PEMELIHARAAN | SPESIFIKASI KEMASAN
1 | CB3012 Plus | Dirancang untuk digunakan sebagai tempat tidur pasien rawat inap pada lembaga medis, seperti rumah sakit, puskesmas, klinik dsb, juga digunakan untuk fasilitas perawatan orang tua di lembaga panti jompo | untuk pemasangan H/F Board pada Main frame Nursing Bed adalah dengan cara mengaitkan Bracket hook yang terdapat di H/F Board pada Pin Stopper, lalu pastikan terkunci dengan stopper lock | 1. Ranjang harus dibersihkan dan didesinfeksi setiap minggu dan atau sebelum pasien baru menggunakan ranjang 2. Matrass harus dilepaskan dari ranjang untuk dibersihkan | CB/F K200/M150x3/K200 2050 x 1070 mm x 190
|  | Digunakan untuk pasien dewasa dan tidak diperuntukkan untuk pasien bayi atau anak-anak | pemasangan matrass dilakukan secara manual dengan cara meletakkan matras pada permukaan ranjang, matras yang di gunakan matras dengan ukuran dan jenis yang benar sesuai rekomendasi pabrik. Matras yang tidak kompatibel bisa menimbulkan bahaya. | Setelah ranjang dibersihkan seka semua permukaan ranjang dengan natrium dikloroisosianuat (NADCC) pada konsentrasi 1000 bagian per juta (0.1%) dari klorin yang tersedia
|  | Tempat tidur tidak boleh digunakan untuk dua orang atau lebih, karena akan melebihi beban maksimum dari tempat tidur | Pada saat penggunaan crank handle maka crank handle dikeluarkan dari tempatnya dan setelah penggunaan crank handle dimasukkan kembali ke tempatnya. cara kerja CB3012 Plus terdapat 2 Crank yang memiliki fungsi tersendiri terkait pengaturan Posisi Bottom Frame. Ketika handle di gerakkan searah jarum jam, maka Crank akan memutar dan mengontrol Pipa besi Torsi. Gerakkan Handle Crank 3 (paling kanan) searah jarum jam, dan Head rest akan bergerak, Gerakkan Handle Crank 2 (tengah) searah jarum jam, dan Lower Frame akan bergerak naik dan turun | crank handle harus di cek secara berkala dengan cara diputar kekiri dan kekanan serta ditambahkan grease jika sudah tidak normal
|  | Ranjang ini di bangun dengan karakteristik kontruksi besi pipa dan pelat meliputi permukaan berbaring, Kontruksi permukaan bawah dan aksesoris. Untuk perawatan yang lebih ergonomis, Nursing Bed CB 3012 Series bekerja secara manual menggunakan sistem Crank, untuk menyesuaikan ketinggian ranjang (High-Low) dan pada posisi duduk yang nyaman (Head Rest) | Penguncian castor dilakukan secara manual dengan cara menekan lock handle pada caster dengan menggunakan kaki | Periksa kelancaran dalam pemasangan dan pelepasan side rail
|  | Penggunaan double lock caster (satu di depan dan satu dibelakang) berfungsi untuk memberikan kestabilan dan keamanan pada saat posisi ranjang dalam kondisi terparkir | side rail sebagai pengaman di sisi kiri dan kanan ranjang dipasang sesuai dengan lubang yang tersedia. Langkah pemasangan Side Rail adalah sebagai berikut : - Siapkan Side Rail dan Posisikan tepat pada insert yang tersedia pada Side Frame, Pada saat akan instalasi, Posisikan Side Rail pada posisi di atas - Pasang Bolt beserta Ring pada Insert yang telah tersedia, dan kencangkan menggunakan Kunci L (L key) atau Screwdriver (+) sampai dengan batas pengencangan bolt 4. Langkah melipat dan menaikkan Side Rail adalah sebagai berikut - Tarik kunci side rail yang berada pada frame side rail - Lepas tarikan kunci dan biarkan Side Rail bergerak sendiri secara perlahan ke arah bawah | lakukan pemeriksaan visual kondisi castor secara rutin untuk memastikan castor masih berfungsi secara baik
|  | Ranjang digunakan pada posisi yang bisa berpindah pindah karena ranjang dilengkapi dengan roda castor yang berfungsi untuk memudahkan saat dipindah |  | periksa matrass jika terjadi kerusakan dan kemungkinan terjadinya air masuk kedalam matrass

## CB3012Plus petunjuk (eng) (hidden)
INSTRUCTIONS FOR USE, TRAINING MATERIALS, INSTALLATION AND MAINTENANCE INSTRUCTIONS
No | TYPE | USE | INSTALLATION | MAINTENANCE | PACKAGING SPECIFICATIONS
1 | CB3012 Plus | Designed for use as inpatient beds in medical institutions, such as hospitals, health centers, clinics, etc., also used for elderly care facilities in nursing home institutions | To install the H/F Board on the Main frame of the Nursing Bed, you need to attach the Bracket hook on the H/F Board to the Pin Stopper, then make sure it is locked with the stopper lock. | 1. The bed must be cleaned and disinfected every week and/or before a new patient uses the bed 2. The mattress must be removed from the bed for cleaning | CB/F K200/M150x3/K200 2050 x 1070 mm x 190
|  | Used for adult patients and not intended for infants or children | Installation of the mattress is done manually by placing the mattress on the surface of the bed, using a mattress of the correct size and type according to the manufacturer's recommendations. Incompatible mattresses can be dangerous. | After the bed is cleaned wipe all bed surfaces with sodium dichloroisocyanuate (NADCC) at a concentration of 1000 parts per million (0.1%) of available chlorine
|  | The bed should not be used for two or more people, as this will exceed the maximum load of the bed | When using the crank handle, the crank handle is removed from its place and after use the crank handle is put back in its place. How the CB3012 Plus works, there are 2 cranks which have their own functions related to setting the bottom frame position. When the handle is moved clockwise, the crank will rotate and control the iron pipe torque. Move Crank Handle 3 (far right) clockwise, and the Head rest will move, Move Crank Handle 2 (middle) clockwise, and the Lower Frame will move up and down | The crank handle must be checked periodically by turning it left and right and adding grease if it is not normal
|  | This bed is built with the characteristics of iron pipe and plate construction including a lying surface, bottom surface construction and accessories. For more ergonomic care, the CB 3012 Series Nursing Bed works manually using a Crank system, to adjust the height of the bed (High-Low) and a comfortable sitting position (Head Rest). | Locking the castor is done manually by pressing the lock handle on the caster using the foot | Check the smooth installation and removal of side rails
|  | The use of double lock casters (one at the front and one at the back) functions to provide stability and security when the bed is parked. | Side rails as safety on the left and right sides of the bed are installed according to the available holes. Side Rail installation steps are as follows: - Prepare the Side Rail and position it exactly on the insert provided on the Side Frame. When installing, position the Side Rail in the above position - Install the Bolt and Ring on the Insert provided, and tighten using the L Key or Screwdriver (+) to the bolt tightening limit 4. Steps to fold and raise the Side Rail are as follows - Pull the side rail lock on the side rail frame - Release the key and let the Side Rail slowly move downwards on its own | Perform visual checks on the condition of the castor regularly to ensure the castor is still functioning properly
|  | The bed is used in a position that can be moved because the bed is equipped with castor wheels which function to make it easier to move |  | Check the mattress for damage and the possibility of water entering the mattress

## CB3003Plus petunjuk (hidden)
PETUNJUK PENGGUNAAN, MATERI PELATIHAN, PETUNJUK PEMASANGAN SERTA PEMELIHARAAN
No | TYPE | PENGGUNAAN | PEMASANGAN | PEMELIHARAAN | SPESIFIKASI KEMASAN
2 | CB3003 Plus | Dirancang untuk digunakan sebagai tempat tidur pasien rawat inap pada lembaga medis, seperti rumah sakit, puskesmas, klinik dsb, juga digunakan untuk fasilitas perawatan orang tua di lembaga panti jompo | untuk pemasangan H/F Board pada Main frame Nursing Bed adalah dengan cara mengaitkan Bracket hook yang terdapat di H/F Board pada Pin Stopper, lalu pastikan terkunci dengan stopper lock | 1. Ranjang harus dibersihkan dan didesinfeksi setiap minggu dan atau sebelum pasien baru menggunakan ranjang 2. Matrass harus dilepaskan dari ranjang untuk dibersihkan | CB/F K200/M150x3/K200 2050 x 1070 mm x 190
|  | Digunakan untuk pasien dewasa dan tidak diperuntukkan untuk pasien bayi atau anak-anak | pemasangan matrass dilakukan secara manual dengan cara meletakkan matras pada permukaan ranjang, matras yang di gunakan matras dengan ukuran dan jenis yang benar sesuai rekomendasi pabrik. Matras yang tidak kompatibel bisa menimbulkan bahaya. | Setelah ranjang dibersihkan seka semua permukaan ranjang dengan natrium dikloroisosianuat (NADCC) pada konsentrasi 1000 bagian per juta (0.1%) dari klorin yang tersedia
|  | Tempat tidur tidak boleh digunakan untuk dua orang atau lebih, karena akan melebihi beban maksimum dari tempat tidur | Pada saat penggunaan crank handle maka crank handle dikeluarkan dari tempatnya dan setelah penggunaan crank handle dimasukkan kembali ke tempatnya. cara kerja CB3003 Plus terdapat 3 Crank yang memiliki fungsi tersendiri terkait pengaturan Posisi Bottom Frame. Ketika handle di gerakkan searah jarum jam, maka Crank akan memutar dan mengontrol Pipa besi Torsi. Gerakkan Handle Crank 3 (paling kanan) searah jarum jam, dan Head rest akan bergerak, Gerakkan Handle Crank 2 (tengah) searah jarum jam, dan Lower Frame akan bergerak naik dan turun, Gerakkan Handle Crank 1 (kiri) searah jarum jam, dan Foot Rest akan bergerak naik dan turun | crank handle harus di cek secara berkala dengan cara diputar kekiri dan kekanan serta ditambahkan grease jika sudah tidak normal
|  | Ranjang ini di bangun dengan karakteristik kontruksi besi pipa dan pelat meliputi permukaan berbaring, Kontruksi permukaan bawah dan aksesoris. Untuk perawatan yang lebih ergonomis, Nursing Bed CB 3003 Series bekerja secara manual menggunakan sistem Crank, untuk menyesuaikan ketinggian ranjang (High-Low) dan pada posisi duduk yang nyaman (Head Rest dan Foot Rest) | Penguncian castor dilakukan secara manual dengan cara menekan lock handle pada caster dengan menggunakan kaki | Periksa kelancaran dalam pemasangan dan pelepasan side rail
|  | Penggunaan double lock caster (satu di depan dan satu dibelakang) berfungsi untuk memberikan kestabilan dan keamanan pada saat posisi ranjang dalam kondisi terparkir | side rail sebagai pengaman di sisi kiri dan kanan ranjang dipasang sesuai dengan lubang yang tersedia. Langkah pemasangan Side Rail adalah sebagai berikut : - Siapkan Side Rail dan Posisikan tepat pada insert yang tersedia pada Side Frame, Pada saat akan instalasi, Posisikan Side Rail pada posisi di atas - Pasang Bolt beserta Ring pada Insert yang telah tersedia, dan kencangkan menggunakan Kunci L (L key) atau Screwdriver (+) sampai dengan batas pengencangan bolt 4. Langkah melipat dan menaikkan Side Rail adalah sebagai berikut - Tarik kunci side rail yang berada pada frame side rail - Lepas tarikan kunci dan biarkan Side Rail bergerak sendiri secara perlahan ke arah bawah | lakukan pemeriksaan visual kondisi castor secara rutin untuk memastikan castor masih berfungsi secara baik
|  | Ranjang digunakan pada posisi yang bisa berpindah pindah karena ranjang dilengkapi dengan roda castor yang berfungsi untuk memudahkan saat dipindah |  | periksa matrass jika terjadi kerusakan dan kemungkinan terjadinya air masuk kedalam matrass

## CB3003Plus petunjuk (eng) (hidden)
INSTRUCTIONS FOR USE, TRAINING MATERIALS, INSTALLATION AND MAINTENANCE INSTRUCTIONS
No | TYPE | USE | INSTALLATION | MAINTENANCE | PACKAGING SPECIFICATIONS
2 | CB3003 Plus | Designed for use as inpatient beds in medical institutions, such as hospitals, health centers, clinics, etc., also used for elderly care facilities in nursing home institutions | To install the H/F Board on the Main frame of the Nursing Bed, you need to attach the Bracket hook on the H/F Board to the Pin Stopper, then make sure it is locked with the stopper lock. | 1. The bed must be cleaned and disinfected every week and/or before a new patient uses the bed 2. The mattress must be removed from the bed for cleaning | CB/F K200/M150x3/K200 2050 x 1070 mm x 190
|  | Used for adult patients and not intended for infants or children | Installation of the mattress is done manually by placing the mattress on the surface of the bed, using a mattress of the correct size and type according to the manufacturer's recommendations. Incompatible mattresses can be dangerous. | After the bed is cleaned wipe all bed surfaces with sodium dichloroisocyanuate (NADCC) at a concentration of 1000 parts per million (0.1%) of available chlorine
|  | The bed should not be used for two or more people, as this will exceed the maximum load of the bed | When using the crank handle, the crank handle is removed from its place and after use the crank handle is put back in its place. How the CB3003 Plus works, there are 3 cranks which have their own functions related to setting the bottom frame position. When the handle is moved clockwise, the crank will rotate and control the iron pipe torque. Move Crank Handle 3 (far right) clockwise, and the Head rest will move, Move Crank Handle 2 (middle) clockwise, and the Lower Frame will move up and down, Move Crank Handle 1 (left) clockwise, and The Foot Rest will move up and down | The crank handle must be checked periodically by turning it left and right and adding grease if it is not normal
|  | This bed is built with the characteristics of iron pipe and plate construction including a lying surface, bottom surface construction and accessories. For more ergonomic care, the CB 3003 Series Nursing Bed works manually using a Crank system, to adjust the height of the bed (High-Low) and a comfortable sitting position (Head Rest and Foot Rest) | Locking the castor is done manually by pressing the lock handle on the caster using the foot | Check the smooth installation and removal of side rails
|  | The use of double lock casters (one at the front and one at the back) functions to provide stability and security when the bed is parked. | Side rails as safety on the left and right sides of the bed are installed according to the available holes. Side Rail installation steps are as follows: - Prepare the Side Rail and position it exactly on the insert provided on the Side Frame. When installing, position the Side Rail in the above position - Install the Bolt and Ring on the Insert provided, and tighten using the L Key or Screwdriver (+) to the bolt tightening limit 4. Steps to fold and raise the Side Rail are as follows - Pull the side rail lock on the side rail frame - Release the key and let the Side Rail slowly move downwards on its own | Perform visual checks on the condition of the castor regularly to ensure the castor is still functioning properly
|  | The bed is used in a position that can be moved because the bed is equipped with castor wheels which function to make it easier to move |  | Check the mattress for damage and the possibility of water entering the mattress

## CB3300Plus petunjuk (hidden)
PETUNJUK PENGGUNAAN, MATERI PELATIHAN, PETUNJUK PEMASANGAN SERTA PEMELIHARAAN
No | TYPE | PENGGUNAAN | PEMASANGAN | PEMELIHARAAN | SPESIFIKASI KEMASAN
3 | CB3300 Plus | Dirancang untuk digunakan sebagai tempat tidur pasien rawat inap pada lembaga medis, seperti rumah sakit, puskesmas, klinik dsb, juga digunakan untuk fasilitas perawatan orang tua di lembaga panti jompo | untuk pemasangan H/F Board pada Main frame Nursing Bed adalah dengan cara mengaitkan Bracket hook yang terdapat di H/F Board pada Pin Stopper, lalu pastikan terkunci dengan stopper lock | 1. Ranjang harus dibersihkan dan didesinfeksi setiap minggu dan atau sebelum pasien baru menggunakan ranjang 2. Matrass harus dilepaskan dari ranjang untuk dibersihkan | CB/F K200/M150x3/K200 2050 x 1070 mm x 190
|  | Digunakan untuk pasien dewasa dan tidak diperuntukkan untuk pasien bayi atau anak-anak | pemasangan matrass dilakukan secara manual dengan cara meletakkan matras pada permukaan ranjang, matras yang di gunakan matras dengan ukuran dan jenis yang benar sesuai rekomendasi pabrik. Matras yang tidak kompatibel bisa menimbulkan bahaya. | Setelah ranjang dibersihkan seka semua permukaan ranjang dengan natrium dikloroisosianuat (NADCC) pada konsentrasi 1000 bagian per juta (0.1%) dari klorin yang tersedia
|  | Tempat tidur tidak boleh digunakan untuk dua orang atau lebih, karena akan melebihi beban maksimum dari tempat tidur | Nursing Bed CB 3300 series di bangun dengan karakteristik kontruksi besi pipa dan pelat meliputi permukaan berbaring, Kontruksi permukaan bawah dan aksesoris. Untuk perawatan yang lebih ergonomis, Nursing Bed CB 3300 series bekerja secara manual menggunakan sistem elektrik menggunakan Motor, untuk menyesuaikan ketinggian ranjang (High-Low, Foot Rest, Head Rest) pada posisi duduk yang nyaman dan posisi kejut atau posisi jantung.CB330 Plus menggunakan Penggerak motorik sistem Single Actuator yang di aplikasikan pada pipa penggerak Platform Matras untuk menggerakkan sandaran punggung dan kaki pada Bed CB 3300 | 1. Lepaskan sambungan listrik ranjang dari pasokan listrik sebelum memulai kegiatan pembersihan atau pemeliharaan. 2. Untuk menghindari Dekontaminasi, Jangan biarkan steker listrik atau kabel catu daya basah. 3. Jika handheld jatuh ke permukaan yang keras, periksa apakah semua tombol berfungsi dengan benar. 4. Letakkan handheld pada rel sisi dengan menggunakan klip pada bagian belakangnya; hal ini akan membantu mencegah pengoperasian kendali yang tidak disengaja.
|  | Ranjang ini di bangun dengan karakteristik kontruksi besi pipa dan pelat meliputi permukaan berbaring, Kontruksi permukaan bawah dan aksesoris. Untuk perawatan yang lebih ergonomis, Nursing Bed CB 3300 Plus bekerja secara otomatis dengan sistem Motor Actuator yang dikendalikan dari handheld Controller | Penguncian castor dilakukan secara manual dengan cara menekan lever Pipe dengan menggunakan kaki | Periksa kelancaran dalam pemasangan dan pelepasan side rail
|  | Penggunaan Total lock caster berfungsi untuk memberikan kestabilan dan keamanan pada saat posisi ranjang dalam kondisi terparkir | side rail sebagai pengaman di sisi kiri dan kanan ranjang dipasang sesuai dengan lubang yang tersedia. Langkah pemasangan Side Rail adalah sebagai berikut : - Siapkan Side Rail dan Posisikan tepat pada insert yang tersedia pada Side Frame, Pada saat akan instalasi, Posisikan Side Rail pada posisi di atas - Pasang Bolt beserta Ring pada Insert yang telah tersedia, dan kencangkan menggunakan Kunci L (L key) atau Screwdriver (+) sampai dengan batas pengencangan bolt 4. Langkah melipat dan menaikkan Side Rail adalah sebagai berikut - Tarik kunci side rail yang berada pada frame side rail - Lepas tarikan kunci dan biarkan Side Rail bergerak sendiri secara perlahan ke arah bawah | lakukan pemeriksaan visual kondisi castor secara rutin untuk memastikan castor masih berfungsi secara baik
|  | Ranjang digunakan pada posisi yang bisa berpindah pindah karena ranjang dilengkapi dengan roda castor yang berfungsi untuk memudahkan saat dipindah |  | periksa matrass jika terjadi kerusakan dan kemungkinan terjadinya air masuk kedalam matrass

## CB3300Plus petunjuk (eng)
INSTRUCTIONS FOR USE, TRAINING MATERIALS, INSTALLATION AND MAINTENANCE INSTRUCTIONS
No | TYPE | USE | INSTALLATION | MAINTENANCE
3 | CB3300 Plus | Designed for use as inpatient beds in medical institutions, such as hospitals, health centers, clinics, etc., also used for elderly care facilities in nursing home institutions | To install the H/F Board on the Main frame of the Nursing Bed, you need to attach the Bracket hook on the H/F Board to the Pin Stopper, then make sure it is locked with the stopper lock. | 1. The bed must be cleaned and disinfected every week and/or before a new patient uses the bed 2. The mattress must be removed from the bed for cleaning
|  | Used for adult patients and not intended for infants or children | Installation of the mattress is done manually by placing the mattress on the surface of the bed, using a mattress of the correct size and type according to the manufacturer's recommendations. Incompatible mattresses can be dangerous. | After the bed is cleaned wipe all bed surfaces with sodium dichloroisocyanuate (NADCC) at a concentration of 1000 parts per million (0.1%) of available chlorine
|  | The bed should not be used for two or more people, as this will exceed the maximum load of the bed | The CB 3300 series nursing bed is built with the characteristics of iron pipe and plate construction including a lying surface, bottom surface construction and accessories. For more ergonomic care, the CB 3300 series Nursing Bed works automaticly using an electrical system using a actuator motor, to adjust the height of the bed (High-Low, Foot Rest, Head Rest) to a comfortable sitting position and a shock position or heart position. The CB 3300 Plus uses Single Actuator system motor drive which is applied to the Mattress Platform drive pipe to move the backrest and legs on the CB 3300 Plus Bed | 1. Disconnect the bed from the power supply before starting any cleaning or maintenance activities. 2. To avoid Decontamination, Do not let the power plug or power supply cable get wet. 3. If the handheld falls onto a hard surface, check whether all buttons work properly Correct. 4. Place the handheld on the side rail using the clip on the back; this thing will helps prevent unintentional control operation.
|  | This bed is built with the characteristics of iron pipe and plate construction including a lying surface, bottom surface construction and accessories. For more ergonomic care, the CB 3300 Plus Nursing Bed works automatically with a Motor Actuator system controlled from a handheld controller | Locking the castor is done manually by pressing the Pipe lever with your foot | Check the smooth installation and removal of side rails
|  | The use of a total lock caster functions to provide stability and security when the bed is parked | Side rails as safety on the left and right sides of the bed are installed according to the available holes. Side Rail installation steps are as follows: - Prepare the Side Rail and position it exactly on the insert provided on the Side Frame. When installing, position the Side Rail in the above position - Install the Bolt and Ring on the Insert provided, and tighten using the L Key or Screwdriver (+) to the bolt tightening limit Steps to fold and raise the Side Rail are as follows - Pull the side rail lock on the side rail frame - Release the key and let the Side Rail slowly move downwards on its own | Perform visual checks on the condition of the castor regularly to ensure the castor is still functioning properly
|  | The bed is used in a position that can be moved because the bed is equipped with castor wheels which function to make it easier to move |  | Check the mattress for damage and the possibility of water entering the mattress