| MAINTENANCE TRAINING MATERIAL FOR BED TYPE CB 7003 D Table 1. Crank Assy Replacement Method and Work Sequence * Hi-Lo Case |  |  |
| --- | --- | --- |
| No | Working Order | Method |
| 1 | Raise the Bed to the highest level | - Raise the Bed to the highest level - Turn the Crank Handle to the right and stop where the upstroke becomes heavier. |
| 2 | Remove both springs | Remove both Springs using the Hook JIG |
| 3 | Lower the Bed to the Lowest level | Turn the Crank Handle to the left and stop where the stoper meets |
| 4 | Remove the Comple Crank Rod End | Release the Snap pin and pull the pin |
| 5 | Remove the Crank Assy | - Open the Instruction Label on the handle and open the screw, then remove the handle. |
|  |  | - Remove the two M10 bolts attached to the cross member. |
|  |  | - Remove the Crank Assy |
| 6 | Installation of Crank Assy | - Install the handle to the cross frame foot side hole and tigh the screw securely |
|  |  | - Install the two M10 bolts and tighten |
|  |  | - Install the end of the rod axis on the torque tube bracket using the pin, collar and snap pin. |
| 7 | Raise the Bed to the highest level | Turn the Crank Handle to the right and stop where the upstroke becomes heavier. |
| 8 | Install both Springs | Install the two springs using the hook JIG |
| 9 | Lower the Bed to the Lowest level | Turn the Crank Handle to the left and stop where the stoper meets |

| MAINTENANCE TRAINING MATERIAL FOR BED TYPE CB 7003 D Table 1. Caster Replacement Method and Work Sequence |  |  |  |  |
| --- | --- | --- | --- | --- |
| No | Urutan Kerja | Metoda/Cara Kerja |  |
| 1 | Prepare everything related to replacement work | Raise the bed by turning the crank handle clockwise until there is enough space (between the floor and casters), then install the stand. |  |
| 2 | Remove the Casters | Loosen the Nut with a 21 mm wrench and remove the caster |  |
| 3 | Install Casters | Insert the caster thread into the foot pipe thread and tighten using a wrench until tigh |  |
| 4 | Make sure the installation is correct | Make sure the 16 diameter Spring Washer is even and has no gaps |  |
| 5 | Take out the chockstand | Turn the crank handle clockwise once there is space then pull the support holder out. |  |
| 6 | Make sure all functions are operational | Make sure the forward and backward rotation is horizontal without any abnormalities. Make sure there are no abnormalities in the lock function. |  |