MAINTENANCE TRAINING MATERIAL FOR BED TYPE  CB 7011 D

Table 1. Crank Assy Replacement Method and Work Order * Hi-Lo Case

| No | Work Order | Method |
| --- | --- | --- |
| 1 | Raise the bed to the highest level | Turn the crank handle to the right and stop where the upstroke become heavier |
| 2 | Remove both Springs | Remove both Springs using the Hook JIG |
| 3 | Lower the Bed to the Lowest level | Turn the Crank Handle to the left and stop where the stoper meets |
| 4 | Remove the Comple Crank Rod End | Release the Snap pin and pull the Pin |
| 5 | Remove the Crank ASSY | 1. Open the Instruction Label on the handle and unscrew the screw, then remove the handle. |
|  |  | 2. Remove the two M10 bolts attached to the cross member |
|  |  | 3. Remove the Crank Assy |
| 6 | Installation of CRANK ASSY | 1. Install the handle to the cross frame foot side hole and tigh the screw securely. |
|  |  | 2. Install the two M10 bolts and tighten |
|  |  | 3. Install the end of the rod axis on the BRKT torque tube using the pin, collar and snap pin. |
| 7 | Raise the Bed to the highest level | Turn the Crank Handle to the right and stop where the upstroke becomes heavier. |
| 8 | Install both Springs | Install the two springs using the hook JIG |
| 9 | Lower the Bed to the Lowest level | Turn the Crank Handle to the left and stop where the stoper meets |

MAINTENANCE TRAINING MATERIAL FOR BED TYPE  CB 7011 D

Table 2. Method and Sequence of Cylinder Replacement Process

| No | Work Order | Method |
| --- | --- | --- |
| Cylinder Removal Process |  |  |
| 1 | Raise the Bed to the highest level | Turn the Crank Handle clockwise and stop where the force /rotation of the crank be comes heavier |
| 2 | Install the Jig (round iron diameter 10 x 8cm) | Insert the Jig between BRKT1 and Cross Frame BRKT while pulling the operating lever |
| 3 | Disconnect Control Cable | Loosen the M5 Nut using a 10 Ring Wrench and remove the Jig. |
|  |  | Loosen the screw with the H+ screwdriver until the cable stopper is loose |
| 4 | Remove the front cylinder from Bottom A | Remove the Snap Pin with pliers and pull the pin out and hang on the BRKT. |
| 5 | Remove the rear cylinder from the BRKT Cross Frame | Remove the Snap Pin with pliers and pull the pin, so that the Cylinder Comple comes off. |
| 6 | Remove the Cylinder from BRKT2, Cylinder cover and Nut. | Turn the Cylinder counterclockwise, first BRKT2 |
| New Cylinder Installation Process |  |  |
| 7 | Install Nut M10 and Cylinder Cover and BRKT2 to Cylinder | Install the Nut, Cylinder Cover and BRKT2 into the Cylinder by inserting and turning it clockwise |
| 8 | Install the Cylinder set to the BRKT Cross Frame | 1. Insert the BRKT2 part into the Cross Frame BRKT after the holes are aligned |
|  |  | 2. Insert the end of the Cylinder into Bottom A BRKT , if the holes are aligned, insert the pin and install the snap pin. |
| 9 | Install the end of the Control Cable | Attach the end of the Control Cable by inserting it through the cable holder and bolt hole on BRKT1. |
| 10 | Install the Jig | Insert the Jig between BRKT1 and Cross Frame BRKT by pressing Bottom A. |

| 11 | Control Cable Settings | The wire on the Control Cable is pulled with pliers and the Nut is tightened with a 10 ring wrench, after that the screw is also tightened with a screwdriver. |
| --- | --- | --- |
| 13 | Operation function check | Operate the lever to ensure that you can raise and lower the bottom without Anomaly." |

MAINTENANCE TRAINING MATERIAL FOR BED TYPE  CB 7011 D

Table 3. Control Cable Replacement Methods and Work Order

| No | Work Order | Method |
| --- | --- | --- |
| Process of Removing Control Cable |  |  |
| 1 | Raise the Bed to the highest level | Turn the Crank Handle clockwise and stop where the force/ rotation of the crank becomes heavier. |
| 2 | Install the Jig (round iron diameter 10x8cm) | Insert the Jig between BRKT1 and Cross Frame BRKT while pulling the operating lever |
| 3 | Disconnect Control Cable | 1. Loosen the M5 Nut using a 10 ring wrench and remove the Jig. |
|  |  | 2. Loosen the screw with the H+ screwdriver until the stopper cable is loose. |

|  |  | 3. Open the clamping plate (2 screws) with the plus driver. |
| --- | --- | --- |
|  |  | 4. Open gap between the Cable Holder Plate and Hook with a min screw driver and remove the Control Cable from the hook. |
| The Process of Installing a New Control Cable |  |  |
| 4 | Install the New Control Cable | 1. Install the pendulum on the Hook, insert it into the cable stopper plate |
|  |  | 2. Install the clamp and insert the new Control Cable then tighten the screw. |
|  |  | 3. Install the JIG between BRKT1 and Cross Frame BRKT by pressing Bottom A |
| 5 | Setting Control Cable | The wire on the Control Cable is pulled using pliers and tightened the Nut using 10 wrench. |
| 6 | Check operation function | Operate the lever to ensure that you can raise and lower the bottom without any anomalies. |

MAINTENANCE TRAINING MATERIAL FOR BED TYPE  CB 7011 D

Table 4. Caster Replacement Methods and Work Order

| No | Work Order | Method |
| --- | --- | --- |
| 1 | Prepare everything related to replacement work | Raise the bed by turning the crank handle clockwise until there is enough space (between the floor and the casters), then install the chockstand. |
| 2 | Remove the Casters | Loosen the Nut with a 21 mm wrench and remove the casters. |
| 3 | Install Casters | Insert the caster thread into the foot pipe thread and tighten using 21 mm wrench until tigh. |
| 4 | Make sure the installation is correct | Make sure the 16 diameter Spring Washer is even, there are no gaps. |
| 5 | Take out the Chockstand | Turn the crank handle clockwise once there is space then pull the support holder out. |
| 6 | Make sure the operation function | Make sure the forward and backward rotation is horizontal without any abnormalities. Make sure there are no abnormalities in the lock function. |