## Table 1
Bed Care Training Type CB 7011 D
Table 1. Crank Assy Replacement Method and Work Order * Hi-Lo Case
No | Work Order | Method
1 | Raise the Bed to the highest level | Turn the Crank Handle to the right and stop where the upstroke becomes heavier
2 | Remove both Springs | Remove both Springs using the Hook JIG
3 | Lower the Bed to the Lowest level | Turn the Crank Handle to the left and stop where the stoper meets
4 | Remove the Comple Crank Rod End | Release the Snap pin and pull the Pin
5 | Remove the Crank ASSY | 1. Open the Instruction Label on the handle and unscrew the screw, then remove the handle.
2. Remove the two M10 bolts attached to the cross member
3. Remove the Crank Assy
6 | Installation of CRANK ASSY | 1. In the Cross Frame Foot Side hole, attach the handle with a screw and tighten
2. Install the two M10 bolts and tighten
3. Install the end of the rod axis on the BRKT torque tube using the pin, collar and snap pin.
7 | Raise the Bed to the highest level | Turn the Crank Handle to the right and stop where the upstroke becomes heavier.
8 | Install both Springs | Install the two springs using the hook JIG
9 | Lower the Bed to the Lowest level | Turn the Crank Handle to the left and stop where the stoper meets

## Table 1 (2)
Bed Care Training Type CB 7011 D
Table 2. Crank Assy Replacement Method and Work Order * Hi-Lo Case
No | Work Order | Method
Cylinder Removal Process
1 | Raise the Bed to the highest level | Turn the Crank Handle clockwise and stop where the force/rotation becomes heavier
2 | Install the Jig (round iron diameter 10 x 8cm) | Insert the Jig between BRKT1 and BRKT Cross Frame while pulling the operating lever
3 | Disconnect Control Cable | Loosen the M5 Nut using a 10 Ring Wrench and remove the Jig.
Loosen the screw with the H+ screwdriver until the cable stopper is loose
5 | Remove the front cylinder from Bottom A | Remove the Snap Pin with pliers and pull the pin out and hang on the BRKT.
6 | Remove the rear cylinder from the BRKT Cross Frame | Remove the Snap Pin with pliers and pull the pin, so that the Cylinder Comple comes off.
7 | Remove the Cylinder from BRKT2, Cylinder cover and Nut. | Turn the Cylinder counterclockwise, first BRKT2
Proses pemasangan Cylinder yang baru
8 | Install Nut M10 and Cylinder Cover and BRKT2 to Cylinder | Install the Nut, Cylinder Cover and BRKT2 into the Cylinder by inserting it and turning it clockwise
9 | Install the Cylinder set to the BRKT Cross Frame | 1. Insert the BRKT2 part into the BRKT Cross Frame after the holes are aligned
2. Insert the end of the Cylinder into BRKT Bottom A, if the holes are aligned, insert the pin and install the snap pin.
10 | Install the end of the Control Cable | Attach the end of the Control Cable by inserting it through the cable holder and bolt hole on BRKT1.
11 | Install the Jig | Insert the Jig between BRKT1 and BRKT Cross Frame by pressing Bottom A.
12 | Cable Control Settings | The wire on the Cable Control is pulled with pliers and the Nut is tightened with a 10 spanner, after that the screw is also tightened with a screwdriver.
13 | Operation function check | Operate the lever to ensure that you can raise and lower the bottom without Anomaly."

## Table 1 (3) (hidden)
Bed Care Training Type CB 7011 D
Table 2. Crank Assy Replacement Method and Work Order * Hi-Lo Case
No | Work Order | Method
Process of Removing Control Cable
1 | Raise the Bed to the highest level | Turn the Crank Handle clockwise and stop where the force/rotation becomes heavier.
2 | Install the Jig (round iron diameter 10x8cm) | Insert the Jig between BRKT1 and BRKT Cross Frame while pulling the operating lever
3 | Disconnect Control Cable | 1. Loosen the M5 Nut using a 10 spanner and remove the Jig.
2. Loosen the screw with the H+ screwdriver until the cable stopper is loose.
3. Open the clamping plate (2 screws) with the plus driver.
4. Open the Cable Holder Plate and Hook gap with a min screwdriver and remove the Control Cable from the hook.

## Table 1 (4)
Bed Care Training Type CB 7011 D
Table 3. Control Cable Replacement Methods and Work Order
No | Work Order | Method
Process of Removing Control Cable
1 | Raise the Bed to the highest level | Turn the Crank Handle clockwise and stop where the force/rotation becomes heavier.
2 | Install the Jig (round iron diameter 10x8cm) | Insert the Jig between BRKT1 and BRKT Cross Frame while pulling the operating lever
3 | Disconnect Control Cable | 1. Loosen the M5 Nut using a 10 spanner and remove the Jig.
2. Loosen the screw with the H+ screwdriver until the cable stopper is loose.
3. Open the clamping plate (2 screws) with the plus driver.
4. Open the Cable Holder Plate and Hook gap with a min screwdriver and remove the Control Cable from the hook.
Process of Removing Control Cable
4 | Install the New Control Cable | 1. Install the pendulum on the Hook, insert it into the Stopper Cable Plate
2. Install the clamp and insert the new Control Cable then tighten the screw.
3. Install the JIG between BRKT1 and BRKT Cross Frame by pressing Bottom A
5 | Setting Control Cable | The wire on the Control Cable is pulled using pliers and tightened Nur using 10 spanners
6 | Check operation function | Operate the lever to ensure that you can raise and lower the bottom without any anomalies.

## Table 1 (5)
Bed Care Training Type CB 7011 D
Table 4. Caster Replacement Methods and Work Order
No | Work Order | Method
1 | Prepare everything related to replacement work | Raise the bed by turning the crank handle clockwise until there is enough space (between the floor and the caster), then install the support supports.
2 | Remove the Casters | Loosen the Nut with a 21 mm wrench and remove the caster.
3 | Install Casters | The caster thread is inserted into the foot pipe (thread) after it is hard and then tightened using a wrench.
4 | Make sure the installation is correct | Make sure the 16 diameter Spring Washer is flat, there are no gaps.
5 | Take out the shim mount | Turn the crank handle clockwise once there is space then the support holder is pulled out.
6 | Make sure the operation function | Make sure the forward and backward rotation is horizontal without any abnormalities. Make sure there are no abnormalities in the lock function.