Corrective Action Request

CAR ....1.... of ...2.....

Major Minor

| Organization: | PT. ChitoseInternasional |  |  |
| --- | --- | --- | --- |
| Site(s) audited: | State in S2R | Date(s) of audit(s): | 26~27June 2019 |
| Auditor(s): | Tomie Gautama |  |  |
| Standard(s): | ISO 9001:2008 |  |  |
| Organization Representative: | Mr. Agung Tri Wahyu |  |  |
| Area / Department / Process: | QC |  |  |
| Document Ref.: | Schedule Calibration | Standard Ref.: | Clause7.1.5.2 |
| Issue/Rev. Status: | 0 | CAR Close out date: | Next Surveillance |

| Details of Non-Conformity: |  |  |  |
| --- | --- | --- | --- |
| - Lack evidence was found that activity of verification measuring equipment is carried out properly, during audit was found that dual scope Mpor (FISCHER) is not identified to enable status and activity verification is not done as due date that appointed (verification should be done at March 2018) - Statement conformity already clear identified in external certificate calibration, but statement conformity is inappropriate with standard e.g in certificate was noted + 5 micron but in standard noted + 15 micron (certificate number 029/ECA/III/2015 |  |  |  |
| Organization Representative: |  | Auditor: |  |

| Cause Analysis andCorrective Action taken to prevent recurrence: |  |  |  |
| --- | --- | --- | --- |
| Organization Representative: |  | Date: |  |

| Acceptance of Corrective Action / Comments (useadditional sheets if necessary): |  |  |  |
| --- | --- | --- | --- |
| Auditor: |  | Date: |  |

| Response required (in months) | Major |  | Minor |  |
| --- | --- | --- | --- | --- |
| Corrective Action must be addressed within time frame stated. Verification of action will occur at next visit. Additional follow up may be required as indicated. | Define | - Close Out | Define | - Close Out |

Corrective Action Request

CAR ....2.... of ...2.....

Major Minor

| Organization: | PT. ChitoseInternasional |  |  |
| --- | --- | --- | --- |
| Site(s) audited: | State in S2R | Date(s) of audit(s): | 26~27 June 2019 |
| Auditor(s): | Tomie Gautama |  |  |
| Standard(s): | ISO 9001:2008 |  |  |
| Organization Representative: | Mr. Agung Tri Wahyu |  |  |
| Area / Department / Process: | QC |  |  |
| Document Ref.: | Thickness Test Report | Standard Ref.: | Clause 8.5.1 C |
| Issue/Rev. Status: | 0 | CAR Close out date: | Next Surveillance |

| Details of Non-Conformity: |  |  |  |
| --- | --- | --- | --- |
| Although that implementation of activity monitor and measurement is running properly, However the correction and corrective action taken has not yet been done and recorded to ensure conformity of the product, e.g test ketebalan /thickness test date on 15 January 2019, 11 March 2019 and 20 May 2019 was upper than requirement in standard CINT/QC/F-01 that standard data was noted in range (75 + 15) micron actual data was noted 98, 95 and 97 micron. |  |  |  |
| Organization Representative: |  | Auditor: |  |

| Cause Analysis and Corrective Action taken to prevent recurrence: |  |  |  |
| --- | --- | --- | --- |
| Organization Representative: |  | Date: |  |

| Acceptance of Corrective Action / Comments (useadditional sheets if necessary): |  |  |  |
| --- | --- | --- | --- |
| Auditor: |  | Date: |  |

| Response required (in months) | Major |  | Minor |  |
| --- | --- | --- | --- | --- |
| Corrective Action must be addressed within time frame stated. Verification of action will occur at next visit. Additional follow up may be required as indicated. | Define | - Close Out | Define | - Close Out |